## Slide 1
Non-Invasive Skin Phantom Properties Estimation Using Air-Jet Indentation And Active LiDAR Sensing
1
PRESENTED BY
NAPAT NAOWARANGSEE 231512087
Master’s Degree Student, Frontier Engineering
A Hygienic, Repeatable, and Real-Time System for Skin Biomechanics
Notes: “Good afternoon, my name is Napat Naowarangsee, a master's student in Frontier Engineering.
Today I’m excited to present my work on a novel,

## Slide 2
LITERATURE REVIEW
Notes: "First of all, I would like to talk about literature review

## Slide 3
Problem Statement
LITERATURE REVIEW
Notes: let’s first explore the Problem statement that drive this work

## Slide 4
  Clinical Relevance:
    Skin properties are biomarkers for:
      Scleroderma, lymphedema, aging
      Wound healing, pressure ulcer risk
  Challenges in Current Practice:
    Most tools require direct contact
      → Patient discomfort
      → Friction + boundary artifacts
      → Risk of cross-contamination
  Why Measure Skin Mechanics—And Why Non-Contact Matters?
Problem Statement
LITERATURE REVIEW
  Need:
  A non-contact, reliable method to measure in vivo skin properties safely and accurately
Notes: Why Measure Skin Mechanics—And Why Non-Contact Matters?
“Skin mechanics—such as stiffness and elasticity—are important for diagnosing clinical conditions like scleroderma, monitoring lymphedema, and assessing skin aging or wound healing.
However, most current tools require direct skin contact, which can cause patient discomfort, friction-related artifacts, and hygiene concerns in clinical use.
These drawbacks highlight the need for a safe, reliable, and non-contact alternative that can be used in real-time clinical and research settings.”

## Slide 5
  Why Measure Skin Mechanics—And Why Non-Contact Matters?
Problem Statement
LITERATURE REVIEW
Criteria | Contact Methods (Indenters, Suction Devices) | Non-Contact Method (Air-Jet + LiDAR)
Skin Contact | Required (probe or suction chamber) | Not required (air-jet force + depth sensing)
Patient Comfort | Low – may cause discomfort or skin irritation | High – no physical contact
Hygiene & Sterilization | Requires disinfection or disposables | Excellent – no contamination risk
Operator Dependency | High – manual positioning introduces variability | Low – system is automated and repeatable
Cost & Complexity | Moderate to high | Moderate – consumer-grade LiDAR, simple jet
Suitability for Clinics | Limited – workflow bottlenecks, hygiene issues | High – fast, clean, and comfortable
Notes: This brings us to a crucial comparison: how do conventional, contact-based methods stack up against a non-contact approach?
“This table clearly outlines the differences. Traditional methods like indenters or suction devices require direct contact, which introduces several problems: discomfort, sterilization requirements, and operator dependency.
On the other hand, a non-contact method—like air-jet and LiDAR systems—offers precise, automated measurements with no physical contact. It’s cleaner, more consistent, and better suited for real-time clinical use.”

## Slide 6
Problem with
Conventional Techniques
LITERATURE REVIEW
Notes: Let’s take a closer look at those conventional techniques’ technical drawbacks

## Slide 7
  Contact-based tools introduce:
    Friction artifacts
    Operator dependence
    Sterilization/cross-contamination concerns
  Optical Coherence Elastography (OCE):
    High resolution but shallow depth (<2 mm)
    Expensive, complex setup
    Limited portability
  Limitations of Contact-Based and Optical Techniques
Problem with
Conventional Techniques
LITERATURE REVIEW
Notes: “Contact-based systems often suffer from friction artifacts and operator-induced variability, which can distort results—especially in soft or mobile skin areas.
Meanwhile, optical techniques like Optical Coherence Elastography offer impressive resolution, but they only penetrate shallow layers—less than 2 mm. They also require complex alignment and are quite expensive for widespread use.
“That’s why I developed a system that avoids these problems entirely—delivering force without touch.
Let me show you how it works.”

## Slide 8
Overview of Skin
Phantom Properties
LITERATURE REVIEW
Notes: Let’s take a closer look at

## Slide 9
  Skin thickness
  ≈ 0.1 mm → epidermis
  ≈ 1 – 4 mm → dermis
  4 - 10 → mm hypodermis (deep skin/fat layer)
  >27 mm → reaches underlying muscle
  Layers of Human Skin
  (1) https://www.utmb.edu/pedi_ed/CoreV2/Dermatology/page_03.htm?
  (2) https://doi.org/10.1038/s41598-023-44572-y
(2)
(1)
(1)
(1)
Overview of Skin
Phantom Properties
LITERATURE REVIEW
Notes: “Human skin is multilayered—ranging from a thin epidermis to a thicker dermis and a deep skin layer where is a hypodermis layers belong.

## Slide 10
  Skin Phantom
Overview of Skin
Phantom Properties
  T. J. Hall, M. Bilgen, M. F. Insana, and T. A. Krouskop, "Phantom materials for elastography," IEEE Transactions on Ultrasonics, Ferroelectrics, and Frequency Control, vol. 44, no. 6, pp. 1355–1365, Nov. 1997, doi: 10.1109/58.656136.
Material | Pros | Cons
Gelatin (30 – 180 g/L) | Inexpensive and widely available Easy to prepare by simple heating and cooling Tunable stiffness via concentration and bloom strength (4.8 – 158 kPa) | Prone to dehydration and syneresis; requires refrigeration and sealed storage Limited lifespan (≈1 week) due to microbial growth Viscoelastic behavior can vary with temperature and additives
Agarose (5 – 30 g/L) | Heat‐stable hydrogel with predictable gelation Broad stiffness tuning (≈ 7.6 to 195 kPa) by concentration Good optical clarity for imaging applications | Requires precise heating (~95 °C) to dissolve Brittle at low concentrations, prone to cracking Limited lifespan (≈5 days) due to syneresis
  Accessible chemical.
  Cost-effective and easily tunable.
  Limited lifespans and require careful storage.
Relation between Young’s modulus & Concentration (C in g/L)
LITERATURE REVIEW
Notes: In testing, we often simulate these layers using hydrogels like gelatin and agarose, which allow us to tune mechanical properties by adjusting their concentration.”
Gelatin is inexpensive, easy to prepare, and tunable—but it dehydrates easily.
Agarose offers a broader stiffness range and better optical clarity, but it’s brittle at low concentrations and harder to dissolve.
Nevertheless, both materials exhibit tunable Young’s modulus values that closely match biological skin, which makes them ideal for system validation.”

## Slide 11
  Skin Phantom Condition
Overview of Skin
Phantom Properties
Requirement | Specification
Target Young’s modulus (E) | 5 kPa – 190 kPa
Thickness (L) | 20 mm – 30 mm (according to skin thickness)
Surface finish | Matte, diffuse (to optimize LiDAR returns)
Ease of fabrication | Minimal equipment, accessible chemicals
  (1) Mayrovitz, H.N. (2017). Mechanical properties of ballistic gelatin and their relation to tissue-like behavior. J. Biomech. Eng., 139(8):081005.
  (2) Chen, A. et al. (2016). Multilayer tissue‐mimicking phantoms with tunable mechanical properties. Med. Phys., 43(6):3117–3130.
(1)
(2)
LITERATURE REVIEW
Notes: In conclusion of the skin phantom condition

## Slide 12
IMPACT FORCE
Delivering Tunable Forces for Safe, Repeatable Skin Indentation
LITERATURE REVIEW
Notes: But before building the physical system, I needed to make sure the forces it applies would fall within the safe and meaningful range for skin testing

## Slide 13
  Reference Force range: 0.38 – 14.6
  In vitro macro-indentation forces on phantoms
  Nowak, J., & Kaczmarek, M. K. (2024). Deep Indentation Tests of Soft Materials Using Mobile and Stationary Devices. Materials (Basel, Switzerland), 17(17), 4233. https://doi.org/10.3390/ma17174233
IMPACT FORCE
LITERATURE REVIEW
Notes: I pick up the ref. force range from related work, the in vitro macro-indentation forces on phantoms, which is one candidate closely related to my work, the force range they use to press the phantom is around 0.38 – 14.6 N that will use force my force ref.
So how can we know the impact force without touching

## Slide 14
JET IMPINGEMENT
Notes: “To answer that, I analyzed the physics of jet impingement

## Slide 15
Parameters | Formula / Concept | Notes
Potential-core length | ≈ 6d | Empirical for turbulent jets
Exit velocity | = | Choked, isentropic, 𝛾 = 1.4
Centerline velocity |  | Decay beyond
Jet half-angle 𝛼 | ≈ 10°–12° in established region | Turbulent spread
Impingement diameter ​ |  | Circular cross-section
Impact area 𝐴 | π | Reaching ≈ 0 Pa at 𝑟 = /2
Exit density | = | 
Centerline overpressure ΔP |  | Dynamic head on surface
Pressure distribution P(𝑟) | /( | Parabolic decay from center
Impact force at the center |  | 
Parameters | Formula / Concept | Notes
 |  | Empirical for turbulent jets
 |  | Choked, isentropic, 𝛾 = 1.4
Jet half-angle 𝛼 | ≈ 10°–12° in established region | Turbulent spread
 |  | Circular cross-section
Impact area 𝐴 |  | 
Centerline overpressure ΔP |  | Dynamic head on surface
Pressure distribution P(𝑟) |  | Parabolic decay from center
Impact force at the center |  | 
JET IMPINGEMENT
Notes: —how the airflow from the nozzle spreads, decays, and applies force to the tissue surface.”
“This part of the analysis models the jet behavior using core fluid mechanics principles. We start by estimating the potential core length and exit velocity under choked flow. From there, we compute the decay in centerline velocity and the jet’s half-angle expansion. This helps us estimate the impingement diameter—the width of the contact area—and ultimately, the impact pressure distribution across the surface.”
“The key outcome is the ability to calculate dynamic overpressure at the center and model how that pressure decreases radially. This parabolic pressure profile tells us both the maximum force applied and how it spreads across the region of interest. These insights are critical for ensuring that the system delivers a controlled, uniform force to the phantom during testing.”

## Slide 16
LiDAR Camera
LITERATURE REVIEW
Light detection and ranging
Notes: Next the main equipment in our work

## Slide 17
LiDAR Camera
Intel® RealSense™ LiDAR Camera L515
  Camera Assembly
LITERATURE REVIEW

## Slide 18
LiDAR Camera
Intel® RealSense™ LiDAR Camera L515
  Camera Definition
Feature | Specification
Technology | LiDAR (Time-of-Flight)
Depth Resolution | Up to 1024 × 768
Frame Rate | Up to 30 fps
Depth Accuracy | < 5 mm at 1 m distance
Depth Distance | ~0.49 m – 3.5 m
Wavelength | 860 nm (infrared laser)
RGB Sensor | 1920 × 1080 @ 30 fps
Field of View (FOV) | H: 70° × V: 55° (depth)
    LiDAR operates by emitting laser pulses and measuring the time it takes for the light to return after hitting an object, allowing for the creation of detailed 3D maps.
LiDAR Working Principle
LITERATURE REVIEW
Notes: “The Intel RealSense L515 is a compact LiDAR-based depth sensor that uses time-of-flight technology. It emits near-infrared laser pulses and measures how long they take to bounce back from the surface—creating a dense 3D map of the scene.”

## Slide 19
OBJECTIVE
Notes: So from all I’m saying let’s warp up again in the work objective

## Slide 20
OBJECTIVE
🔸 Design and integrate an air-jet module capable of delivering controlled forces with tunable pressure and nozzle configurations.
🔸 Implement a high-resolution depth tracking system using Intel RealSense L515 to capture sub-millimeter surface deformation in real-time.
🔸 Validate the system on phantoms by evaluating force repeatability, surface displacement accuracy, and computed stiffness / Young’s modulus within safe biomechanical thresholds.
Notes: “The objective of this project was threefold:
First, to design and implement an air-jet module capable of delivering controlled, repeatable forces using adjustable pressure and nozzle configurations.
Second, to integrate a high-resolution LiDAR system—specifically the Intel RealSense L515—to capture sub-millimeter surface deformation in real time.
And third, to validate the system using soft-tissue phantoms by evaluating force accuracy, deformation tracking, and mechanical property estimation such as stiffness and Young’s modulus. Ultimately, the goal is to create a non-invasive, reliable alternative to traditional skin mechanics measurement.”

## Slide 21
METHODOLOGY
Notes: "First of all, let's explore the problem statement that drives this work."

## Slide 22
Parameter | Value (SI)
Potential-core length | 6 × 0.002=0.012 m
Exit velocity | 317 m/s
Centerline velocity | 317⋅(0.012/0.50) ≈ 7.6 m/s
Jet half-angle 𝛼 | 11° (assumed)
Impingement diameter ​ | 0.002+2⋅0.50⋅tan11° ≈ 0.196 m
Impact area 𝐴 | 0.0303 m² (≈ 302.89 cm²)
Centerline overpressure Δ | Pa
Pressure distribution (𝑟) | Pa
Impact force at the center | 5.807 N
Parameter | Value (SI)
 | 6 × 0.002=0.012 m
 | 317 m/s
Jet half-angle 𝛼 | 11° (assumed)
 | 0.002+2⋅0.50⋅tan11° ≈ 0.196 m
Impact area 𝐴 | 0.0303 m² (≈ 302.89 cm²)
Impact force at the center | 5.807 N
  Force range: 0.38 – 14.6 N
METHODOLOGY
    Condition
    Pressure from the air blower 0.9 MPa
    Distance between camera and phantom: 50 cm
    Nozzle diameter 2 mm
    Ideal gas with choked flow

## Slide 23
ROI CALCULATION
(Region of Interest)
METHODOLOGY

## Slide 24
Mechanical Properties | 
Phantom material | E (Young’s modulus) kPa
Gelatin Phantom (30 – 180 g/L) | 4.8 – 158
Agarose Phantom (5 – 30 g/L) | 7.6 - 195
From Young’s modulus;
Calculate ROI selection (A) by;
  Phantom thickness 30 mm
ROI CALCULATION
  Skin Phantom Condition
  Force : 5.807 N
METHODOLOGY
  L from LiDAR

## Slide 25
Next Steps and
Development Pathway
Notes: "First of all, let's explore the problem statement that drives this work."

## Slide 26
Next Steps and
Development Pathway
✅ Current Progress
  📌 Identified suitable phantom materials and concentrations (e.g., gelatin, agarose)
  🎯 Defined Region of Interest (ROI) for LiDAR measurement
  💨 Determined air-blower specifications for force delivery (1–15 N)
  🛠️ Completed draft of experimental setup combining air-jet and LiDAR
  Experiment Set Up
  Distance 50 cm

## Slide 27
Next Steps and
Development Pathway
Phantom Fabrication
Produce phantoms with known mechanical properties and controlled thickness
Ensure reproducibility for stiffness calibration
Force Calibration and Jet Characterization
Quantify force vs. pressure relation at nozzle exit
Validate jet impact force distribution over ROI
System Integration and Testing
Combine air-jet and LiDAR in controlled lab conditions
Run pilot tests to verify repeatable deformation tracking
Data Analysis and Model Fitting
Calculate stiffness, Young’s modulus from deformation data
Compare with known phantom properties for validation
🔄 Next Steps

## Slide 28
THANK YOU
Q & A
Notes: Thank you for your attention

## Slide 29
  Establish whether the flow is choked at the nozzle throat
  For ideal gases with γ = 1.4 (air), choked (sonic) flow occurs whenever the nozzle‐entry (stagnation) to exit (ambient) pressure ratio exceeds:
    Here, / = 1 MPa/0.1 MPa = 10 ≫1.893, so the exit Mach number is ​ =1 (choked flow)
  Flow Behavior
= ​
  https://en.wikipedia.org/wiki/Choked_flow
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 30
  Establish whether the flow is choked at the nozzle throat
  From isentropic relations, the static pressure at the throat/exit is:
  The total (stagnation) density is given by the ideal-gas law at , ​:
=
  https://en.wikipedia.org/wiki/Choked_flow
PARAMETERS CALCULATION
  Air Density
JET IMPINGEMENT

## Slide 31
  Establish whether the flow is choked at the nozzle throat
  Combining the two gives
  For air (𝛾=1.4); = 0.634
  So,
  Air Density
=
= 7.40 kg/
  https://en.wikipedia.org/wiki/Choked_flow
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 32
  Compute the nozzle‐exit velocity
    Under choked, isentropic conditions, the exit static temperature is
    and the sonic speed at the exit is
      Thus, the jet issues at about 316 m/s
= ​
=​
  https://en.wikipedia.org/wiki/Choked_flow
PARAMETERS CALCULATION
  Velocity
JET IMPINGEMENT

## Slide 33
  Centerline velocity decay of an axisymmetric free jet
  For 𝑥 beyond the potential‐core length ​ , the centerline velocity decays approximately as
  where ​ is the axial distance at which / first falls to unity (end of potential core)
    Empirically, for many nozzles ≈ 6 𝑑 (Turbulent Flow).
    If the nozzle diameter is 𝑑=2 mm, then ≈ 0.012 m.
  At x = 0.50 m,
  Velocity
,​
= ​
  https://ntrs.nasa.gov/api/citations/19720016717/downloads/19720016717.pdf
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 34
  Dynamic (stagnation) pressure at the skin
  When a jet is brought to rest on an object (stagnated), the local pressure rise Δ is the dynamic pressure:
  Stagnation Pressure
= * 213.71 Pa,​
  https://en.wikipedia.org/wiki/Stagnation_pressure
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 35
  Stagnation pressure at 𝑟 = 0:
  Radial decay:
  Key takeaway:
    Although the blower supplies 1 MPa, at 0.5 m the skin “sees” only ≈ 213.71 Pa at the center,
    over an area of ≈ 300 cm², with a smooth parabolic pressure decay towards the edges.
  Pressure Distribution P(𝑟)
213.71 Pa,​
​
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 36
  Integral expression for force
  For an axisymmetric pressure distribution (𝑟) (gauge over ambient) on a flat circular spot, the total force is
  where radius at which (𝑟) first returns to zero.
PARAMETERS CALCULATION
  Impact Force
JET IMPINGEMENT

## Slide 37
  Parabolic (stagnation‑point) model
  We previously showed near the centerline
  with
  Thus, the overpressure falls roughly as , reaching ≈ 0 Pa at 𝑟 = /2
​ ,
PARAMETERS CALCULATION
  Impact Force
JET IMPINGEMENT

## Slide 38
  Closed‑form integral
  Plugging into the integral
  Thus, the total force depends only on the peak overpressure ​ , the strain‑rate 𝑘, and the zero‑pressure radius
  = Maximum impact area
  Impact Force
[
[ ]
= ;
PARAMETERS CALCULATION
JET IMPINGEMENT

## Slide 39
Mechanical Properties |  |  |  | 
Phantom material | E (Young’s modulus) kPa | Force (N) | ROI of 5 mm Def. (cm²) | ROI of 10 mm Def. (cm²)
Gelatin Phantom (30 g/L) | 4.16 | 6.452 | 93.053 | 46.526
Gelatin Phantom (180 g/L) | 175 | 6.452 | 2.212 | 1.106
Agarose Phantom (5 g/L) | 7.08 | 6.452 | 54.675 | 27.338
Agarose Phantom (30 g/L) | 201.86 | 6.452 | 1.918 | 0.959
  Impact Area ≈ 303 cm²
5 mm Def. start
10 mm Def. start
5 - 10 mm Def.
The maximum deformation starts at the center and decreases toward the edge
  The ROI was chosen within the RED zone with optimal for LiDAR accuracy
  Phantom thickness 30 mm
  Deformation Range > 5 mm according to LiDAR accuracy
ROI CALCULATION
METHODOLOGY

## Slide 40
  Uniaxial Compression Test
  Overview of Measuring Skin Mechanical Properties
  Mocking skin phantom.
  Measure skin elasticity by applying uniaxial compression
  Non-invasive Skin Properties Estimation Using Air-jet Indentation And Active Lidar Sensing

## Slide 41
  Experiment Set Up
  Minimum Distance 49 cm
  Calculate Skin Mechanical Properties By Using LiDAR

## Slide 42
  Skin Phantom
(3)
Overview of Skin
Phantom Properties
Feature | Gelatin Phantom (10 % w/w) | Agarose Phantom (2 % w/w)
Young’s modulus E | 22 – 40 kPa (10 % gelatin; temperature‐dependent) | ≈100 kPa (2 % agarose, small‐strain)
Tunable range | 10 – 50 kPa by adjusting concentration (5 – 16 % w/w) | 20 – 450 kPa by 1 – 6.6 % w/w range
Surface finish | Naturally matte (good LiDAR reflectance) | Slightly glossy (blot or matte‐spray recommended)
Durability & storage | • Use within 1 week; sealed at 4 °C to avoid drying & mold | • Use within 5 days; sealed & hydrated at 4 °C to prevent syneresis
Pros | – Very low cost; food‐grade ingredients – Easily tunable concentration – No special equipment beyond a refrigerator | – Heat‐stable gel; minimal microbial risk – Wide stiffness tunability – Good long‐term dimensional stability
Cons | – Short lifespan (<1 week); prone to dehydration & spoilage – Mechanical properties drift with temperature | – Requires hot handling and rapid pouring – Brittleness at low concentrations – Limited working window before gelation
LITERATURE REVIEW

## Slide 43
  We know the applied pressure 𝑃 or the force 𝐹 from the air blower.
  LiDAR captures point cloud data in a fixed rectangular ROI before and after indentation.
  We have depth maps (𝑥,𝑦) (before) and (𝑥,𝑦) (after).
  Skin behaves approximately linearly elastic under small deformation.
  Calculate Skin Mechanical Properties By Using LiDAR
  Non-invasive Skin Properties Estimation Using Air-jet Indentation And Active Lidar Sensing
  Assumptions

## Slide 44
  Step 1: Select Fixed ROI
  Define a fixed window on the LiDAR point cloud (e.g., 5 cm × 5 cm) centered at the indentation site.
  The ROI contains many points: each has (, , ).
  Step-by-Step Method:
  Non-invasive Skin Properties Estimation Using Air-jet Indentation And Active Lidar Sensing
  Step 2: Compute Deformation Map
  For each point 𝑖 in the ROI:
  Where:
  : initial depth (pre-indentation)
  ​ : deformed depth (post-indentation)
  : local indentation depth at point 𝑖
= - ​
  Calculate Skin Mechanical Properties By Using LiDAR

## Slide 45
  Step 3: Compute Global Indentation Parameters
  Maximum Indentation Depth
    This is often used to represent the strain input.
  Average Indentation Depth
    Where 𝑁 is the number of points in the ROI.
  Step-by-Step Method:
  Calculate Skin Mechanical Properties By Using LiDAR
= ​
  Volume of Deformation
    approximate the volume 𝑉 between the two surfaces:
    *
    Where:
    is the area per point,
    e.g., =Δ𝑥⋅Δ𝑦 (depends on resolution).

## Slide 46
  Step 4: Estimate Applied Force
  From air pressure:
      Where:
        : applied pressure (Pa)
        ​ : contact area, estimate from ROI or indentation footprint
  Step-by-Step Method:
  Calculate Skin Mechanical Properties By Using LiDAR
= ​

## Slide 47
  Step 5: Calculate Skin Properties
  Option A: Stress-Strain Formula, if we know skin thickness h:
      Where:
        average deformation from LiDAR
        ​ : area where pressure was applied
  Step-by-Step Method:
  Calculate Skin Mechanical Properties By Using LiDAR
= ​

## Slide 48
  Step 5: Calculate Skin Properties
  Option B: Stiffness Model (if h is unknown), use effective stiffness:
  Step-by-Step Method:
  Calculate Skin Mechanical Properties By Using LiDAR
= ​

## Slide 49
  : mean deformation (from LiDAR)
  ​ : peak indentation depth
  𝑉: volume of indentation
  𝐸: estimated Young’s modulus
  𝑘: effective stiffness
  Final Output:
  Calculate Skin Mechanical Properties By Using LiDAR

## Slide 50
  LiDAR L515 accuracy
    Less than 5 mm at a distance of 1 meter and less than 14 mm at 9 meters, under optimal conditions with 95% reflectivity.
    The appropriate deformation that LiDAR is likely to detect is 5 – 10 mm at a distance of 1 meter.
ROI CALCULATION

## Slide 51
  LiDAR CALIBRATION
  Calibration Step
  Set the object away from LiDAR 50 cm
  Measuring in 2 cm x 2 cm ROI
  Measuring 60 times at the same distance (60s)
  Calculate accuracy(Bias) and precision(Uncertainty)
  Increase distance by 10 cm at a time to 60 cm
  Plot graph

## Slide 52
  LiDAR CALIBRATION
  Results
Reference Distance (mm) | Bias (mm) | Uncertainty (Std Dev) (mm) | Mean Measured Distance (mm)
500 | 0.11 | 0.7 | 500.11
510 | -0.65 | 0.78 | 509.35
520 | -0.14 | 0.82 | 519.86
530 | -0.23 | 1 | 529.77
540 | 0.38 | 0.95 | 540.38
550 | 0.02 | 0.97 | 550.02
560 | -0.08 | 0.89 | 559.92
570 | 1.69 | 1.15 | 571.69
580 | -2.03 | 1.33 | 577.97
590 | -0.25 | 1.16 | 589.75
600 | 0.59 | 0.99 | 600.59
  500.11 ± 1.36 mm; [498.75 mm, 501.47 mm].
  550.02 ± 1.90 mm; [548.12 mm, 551.92 mm].