## Slide 1
Presentation XI
Thu Ta Zaw
2415121095
  Literature Review

## Slide 2
  Monte Carlo simulation is used for simulating light transport in biological tissue.
  Integrating spheres make experimental measurements of the reflectance and transmittance of a sample.
  In this research, the Integrating Sphere Monte Carlo (ISMC), which is the extension of Monte Carlo Photon transport methods to simulate the propagation of photons in Integrating Sphere is presented.
  Validated with previous work on Monte Carlo photon transport, and experimental measurements of a synthetic tissue phantom validate this method.
  Presented tow approach lookup table method and multivariable minimization method to back-calculate the optical properties of samples.
  Prediction of tissue optical properties using the Monte Carlo modeling of photon transport in turbid media and integrating sphere

## Slide 3
  Solved problem of standard Monte Carlo (MC) models ignore integrating sphere geometry.
  Integrating spheres cause photon re-incidence and sphere-to-sphere crosstalk, which can lead to inaccurate inverse solution with error over >20%.
  ISMC models photon interactions with integrating spheres and accounts for photon re-incidence and dual-sphere crosstalk.
  It is computationally efficient and avoids full ray-tracing inside spheres and only requires only one simulation per illumination.
  Supports both single and dual sphere configurations.
  Experimental and simulation uncertainties are propagated through both methods.
  Prediction of tissue optical properties using the Monte Carlo modeling of photon transport in turbid media and integrating sphere

## Slide 4
  Integrating Sphere measurement (Reflectance)
  Integrating sphere measurements determine two important values: the measured reflectance, and the Measured transmittance, .
  Calculating requires one value and three measurements:
        : the reflectance of a reflectance standard.
        R( , ) : the measured reflected power of the reflectance standard.
        R(, ) : the measured reflected power of the sample.
        R(0,0) : the measured reflected power of the empty integrating sphere.
  , is provided by the manufacture of the standard. (mostly used the one with the slightly higher reflectance).
  R(0,0) is not identically 0,due to diffraction and laser is not perfectly collimated beam, and thus some light may still reflect.

## Slide 5
  Integrating Sphere measurement (Reflectance)
Fig. Configurations necessary to make the measurements required for the calculation of
Notes: 𝑟_𝑠𝑡𝑑

## Slide 6
  Integrating Sphere measurement (Transmittance)
  Calculating requires the following measurements:
        : the measured transmitted power with the laser blocked or turned off.
        T(, ) : the measured transmitted power of the sample.
        T(0,0) : the measured transmitted power of the empty integrating sphere.
  is not necessarily 0 since the detector in the sphere is not perfect and will likely measure some background even when in darkness.

## Slide 7
  Integrating Sphere measurement (Transmittance)
Fig. Configurations necessary to make the measurements required for the calculation of
Notes: 𝑟_𝑠𝑡𝑑

## Slide 8
  Integrating Sphere measurement (Regular or unscattered transmittance)
  Calculating requires the following measurements:
        : the measured regular transmitted power of the sample.
        : the measured transmitted power with the aperture closed.
        : the measured regular transmitted power of the unabated laser.
  is not always 0 because of external lighting and noise from the detector itself.
  In practice, Mu is difficult to determine accurately due to the massive difference in and

## Slide 9
  Integrating Sphere measurement (Transmittance)
Fig. Configurations necessary to make the measurements required for the calculation of
Notes: 𝑟_𝑠𝑡𝑑

## Slide 10
  Integrating Sphere measurement (Regular or unscattered transmittance)
  The additional required values are:
      The inner wall reflectance
      Diameter of the sphere
      Diameter of each port.
      Maximum permissible polar angle measured from the central axis of the incident laser which will be use as a threshold for regular transmittance.
  The maximum permissible polar angle is and it is calculated from the aperture diameter, A, and the distance of the aperture from the sample, l.

## Slide 11
  Methods
  Extends the standard Monte Carlo photon transport to include integrating sphere.
        Tracks photons after exiting the sample:
        Photons escape through ports
        Detection of photons at the detector
        Absorption of photons by the sphere wall
        Photons re-incidence on sample
  Uses analytical probability for photon re-incidence (no ray tracing)
  Integrating sphere is modeled using the wall reflectance, ports fractional areas and detector fractional areas.
  Accounts for sphere-to-sphere crosstalk which enables using for both single and double integrating sphere.
  Uses Monte Carlo statistics to estimate simulation uncertainty.

## Slide 12
  Assumptions and modifications
  Surface roughness effects is ignored as in common MC simulations sample is placed between the glass slides.
  Baffle is assumed to be large enough to block the photons to directly enters the detector paths and small enough not to block the photons to escapes though the source port.
  Baffle is treated as sphere wall and sphere walls is assumed ideal Lambertian.
  Photons are assumed to exit and re-enter sample at port center, which could cause over estimation of R and T.
  Sample is assumed infinite laterally and side losses are ignored which could overestimate the absorption coefficient.
  For the unbaffled sphere, the photons that directly enter into the detector counted immediately as R or T.
  No charge to re-incidence equations.

## Slide 13
  Verification
  Verification with previous work on Monte Carlo simulation without sphere
  Has Excellent agreement when comparing with the established MC codes for Total reflectance, total transmittance for the index-matched boundary condition.
Table 1. Verification for a sample with a refractive-index-matched boundary

## Slide 14
  Verification
  Verification with previous work on Monte Carlo simulation without sphere
  Achieved excellent agreement for the diffuse reflectance for index-mismatched boundary condition.
Table 2. Verification for a sample with a refractive-index-mismatched boundary

## Slide 15
  Verification
  Verified with the integrating sphere experiment performed on a synthetic tissue phantom
  ISMC simulation results closely match experimental diffuse reflectance and diffuse transmittance.
Table 3. Optical properties of the tissue phantom at wavelength 532 nm according to the National Institute of Standards and Technology (NIST) report

## Slide 16
  Application to measuring optical properties
  Estimate the optical properties, absorption coefficient and reduced scattering by using the diffuse reflectance , diffuse transmittance and optional regular transmittance .
  Lookup Table method
  Simulate the wide range of and values and stores the corresponding R and T results.
  Invert the table to find the optical properties.
  It is easy to understand and suitable for multiple similar samples.
Fig. Lookup table for (a) absorption coefficient and (b) for reduced scattering coefficient
n = 1.4, g = 0.9 and t = 1 cm with a zero-sphere setup.

## Slide 17
  Application to measuring optical properties
  Multivariable minimization method
  Define the objective function for the difference between measured values and simulation result
Notes: Powell’s method don’t require the input function derivative.

## Slide 18
  Propagation of uncertainties
  Lookup table method
  Required both forward propagation and backward propagation.
  Combined experimental and simulation uncertainties for the final uncertainty.
  Minimization method
  Only forward propagation is required.
  Randomly samples the experimental uncertainties (), simulation and sphere parameter uncertainties.
  Run multiple minimizations.
  Final uncertainty is equal to the standard error of the mean.
  Both methods yield realistic and comparable uncertainty estimates.
Notes: Powell’s method don’t require the input function derivative.

## Slide 19
  Propagation of uncertainties
  Lookup table method
Fig. A flowchart showing the propagation of uncertainties through the lookup table method
Notes: Powell’s method don’t require the input function derivative.

## Slide 20
The End
THANK YOU FOR LISTENING
Thu Ta Zaw
2415121095